Spectrometer Live Tool – User Guide

Author: Melih Karaoglan

1. What is this tool?

This application is a live spectrometer interface designed for laboratory measurements and teaching demonstrations. It allows users to visualize spectra in real time, apply reference-based corrections, and export measurement data without modifying the source code.

2. How to start

- Make sure Python is installed.

- Install required libraries:

  - numpy

  - matplotlib

  - seabreeze

- Run the script:

- python spectrometer_tool.py

- If no spectrometer is detected, the tool automatically switches to dummy mode.

3. Interface Overview

The interface consists of:

- A live plot area showing spectra

- Sliders for integration time, smoothing, and ROI

- Checkboxes to show or hide curves

- Buttons for saving data and capturing references

- A status box displaying current settings

4. Typical Lab Workflow

- Start the tool.

- Block the light path and press Dark to record the dark reference.

- Apply the reference condition and press White.

- Adjust integration time and smoothing if needed.

- Observe corrected transmission or absorbance curves.

- Save data manually or enable auto-save.

5. Keyboard Shortcuts

| Key | Action |
| --- | --- |
| d | Capture dark reference |
| w | Capture white reference |
| s | Save current spectrum |
| a | Toggle auto-save |
| p | Pause / resume acquisition |
| o | Choose output folder |
| q | Quit application |

6. Output Data

Saved CSV files contain wavelength and intensity values.
Reference files and measurement data are stored in the selected output folder, allowing easy reuse and post-processing.

7. Notes for Students

- Always capture dark and white references before using corrected or absorbance views.

- Large spikes in corrected data usually indicate poor reference quality.

- Use smoothing and ROI to stabilize peak detection.